## Slide 1: Bookwalter Check Ups

## Slide 2
There Will Always Be One Genuine Bookwalter

## Slide 3: Bookwalter

## Slide 4: Benefits of doing checkups
The checkup is free! Value added selling!
Ensures the Bookwalter® Retractor is in optimal condition
Saves customers’ valuable time
Customized inspection sheet
Samples for repairs
Handling repairs
Creating count sheets
Gaining access to accounts by being a resource not just a sales representative!

## Slide 5
Inspection sheet-full part list
Bookwalter pictures
New products- low profile RoTilt®, lightweight pieces
Samples to sell-malleable blades
Sharpie marker to mark pieces in need of repair/replacement
Identify Inspection Champions:
  Team Leaders/Coordinators: Find out which services use the Bookwalter
Confirm Convenient Time:
  Early morning/late afternoon
  Day of the week (maybe a Saturday is best)
  Confirm Number of Sets and Obtain Count Sheets:
  be sure to ask about other surgical specialties, not just general
Planning Checkups
Items to bring:

## Slide 6
Check List
Basics:
  Remove all any indicators and filters
  Count all instruments to ensure nothing is missing
  Identify true Aspen/Codman vs competitor products
  Make sure count sheets match what is in set (SKUs / descriptions)
  Keep track of what sets you’ve checked and which you have left
Blades & Rachets:
  Check for cracked malleables
  Check for inoperable ratchets
  Ensure there are no rough edges or burrs
Post:
  Check to ensure all table posts and horizontal bars function properly (no loose lips, nothing is bent) –
  Offer lightweight components if applicable
Post Coupling:
  Check to make sure the starbursts match and still tightens
Ring:
  Lay flat and make sure not warped

## Slide 12: Appendix: Check Up Details

## Slide 13: Replace Old or Competitive Table Post with Lightweight Table Post: 50-4688
Different design:
  Hollow shaft
  Reduced dimensions in clamp
  Use of anodized aluminum
Weight loss:
  50-4581 = 4 lbs. 14.4 oz.
  New = 3 lbs. 0.9 oz.
  Total Weight Loss = 1 lbs. 13.5 oz.
38% Reduction in Weight

## Slide 14: Table Post 50-4581
Engage the Table Post clamp by turning the lever clockwise and ensure that it engages all the way down without binding or sticking. If you feel binding or sticking or hear a grinding sound, spray instrument lubricant in the hole in the mounting. Turn the handle all the way in both directions to spread the lubricant.
Reverse the action until the Table Post clamp is fully disengaged. While performing this task, check the mounting bracket top to bottom to ensure there are no gouges or raised spots that will hinder the attachment of the Table Post to the bed rail.
Ensure the cap of the Table Post (on newer models) fits over the top of the Table Post shaft or that the metal square (on older models) easily slides into its resting slot without binding or shifting.
Ensure that the mounting block screws are fully seated and in place. Rub your finger over the screws to make sure they are fully seated.
Make sure the mounting block is not bent or damaged. Run your finger over the edge of the mounting to make sure there are no burrs or rough areas along the edge.
Make sure the handle is not bent.
Ensure attachment pin is not loose and does not have excess play.
Slide the Post Coupling up and down the shaft of the Table Pose to ensure it slides smoothly without binding.
Recommendations:
50-4688 12” Lightweight
50-4697 16” Lightweight
BW0144005 18”
BW0144003 24”

## Slide 15: Identifying Old & Competitive Table Posts
Original Codman® Table Post
Aesculap®
Table
Post

## Slide 16: Identifying Old & Competitive Table Posts
Aesculap® Table Post
Bookwalter®
Heavy Table Post

## Slide 17: Horizontal, Horizontal Flex, and Vertical Extension Bars
Slide the Post Coupling up and down the shaft of the bar to ensure it slides smoothly without binding.
Put the Horizontal Bar on a flat surface and roll the bar to make sure the Horizontal Bar is straight and NOT bent.
Engage the Horizontal Bar clamp and ensure there is no “snag” in the jaw.
Tighten the Horizontal Bar onto the ring and ensure it has a tight grasp without “play” in the ring. Disengage to ensure smooth operation.
For the Horizontal Flex Bar: Turn the lever clockwise to tighten the ball joint mechanism. While holding the shaft, press the distal end of the Flex Bar against a hard surface to make sure that the ball joints don’t slip. Also, ensure that older flex bars have an “anti-backing” screw installed which will prevent the internal ball joint fittings from coming apart.
Recommendations:
50-4552 Horizontal Bar (newer lightweight)
50-4582 Horizontal Flex Bar (newer light weight)
50-4553 Vertical Extension Bar 12” (newer light weight)
50-4797 Vertical Extension bar 18” (new light weight)

## Slide 18: Replace Old or Competitive Horizontal Bar with Lightweight Horizontal Bar: 50-4552
Same product code, but now the shaft is hollow
Weight loss:
  Previous = 2 lbs. 9.4 oz.
  New = 1 lbs. 12.7 oz.
  Total Weight Loss = 12.7 oz.
31% Reduction in Weight
New Design
Old Design with Tongue

## Slide 19: Replace Old or Competitive Horizontal Flex Bar with Lightweight Horizontal Flex Bar: 50-4582
Same product code, but modified design:
  Hollow shaft
  More robust ball-and-socket mechanism and ring attachment
Weight loss:
  Previous = 4 lbs. 6.3 oz.
  New = 2 lbs. 12.9 oz.
  Total Weight Loss = 1 lb. 9.4 oz.
36% Reduction in Weight
New
Design

## Slide 20: Identifying Old Horizontal Flex Bars
Original Codman® Horizontal Flex Bar
Symmetry
Surgical® Horizontal Flex Bar
Narrower Jaw
Wider Jaw
Protruding End on Tightening Mechanism
Flush End on Tightening Mechanism
Narrower Housing for Ball Joints
Wider Housing for Ball Joints

## Slide 21: Post Coupling 50-4554
Slide a Table Post and a Horizontal (Flex) Bar into the two holes of the Post Coupling. Tighten the Post Coupling by turning the lever clockwise. Check the teeth (starburst) on both parts of the Post Coupling to ensure that they are not worn down, gouged, out of line, or blunt. The Post Coupling must easily tighten down on the Table Post and Horizontal (Flex) Bar . (Do not tighten the Post Coupling without having a Table Post and Horizontal (Flex) Bar loaded.)
Check the shaft that connects the teeth. Make sure that it is not bent and that the teeth match up.
Check to make sure that the lever is not bent.
Check to see if the handle stops are removed or broken off.
Check the diameters of the holes to ensure that they are round and not oval from improper tightening. Slide the Horizontal Bar into the holes and make sure there is no binding or gaps in the Post Coupling.
Recommendation:
Newer version of 50-4554 (Lightweight Post Coupling)

## Slide 22: Rings
Check for missing parts on segmented ring (screws, segments, etc.).
Check for cracks at the joints of the segmented ring.
Make sure teeth on the segments line up and are not blunt.
Slide a ratchet around the entire circumference of the ring to ensure it doesn’t stick or bind.
Check for gouges that could tear a surgical glove or hinder movement of a ratchet.
Lay the ring on a flat surface to ensure it is not bent. (If it is a segmented ring, assemble it completely before checking that it is flat.)
Recommendations:
Various sizes of solid and segmented round and oval rings

## Slide 23: Ring Modifications
2001
Spacing of ALL rings changed from 1/4” to 1/8”
This change dramatically improved the functionality of the ratchets on the ring.
Surgeons who have experienced slippage appreciate the importance of this modification. Rotation with no slippage!
Expanded blade placement option 30 degrees.
Surgeon can place retractor blade more off center without the ratchet slipping or skipping along the outside of the ring.

## Slide 24: Identifying Old Rings
Original Codman® Ring
Symmetry Surgical® Ring
Wider Space (1/4”)
Narrower Space (1/8”)

## Slide 25: Rachet Guide

## Slide 26: Straight, Tilt, and RoTilt Ratchets
Slide the ratchet down the post of a blade to ensure the spring is working properly. The ratchet should “click” into each groove on the post. Check the tension of the spring to ensure it is not deformed.
Check the tilting mechanism by depressing the tilt lever and moving the body of the ratchet through the full arc of the tilt to ensure that it smoothly moves without binding into the body of the ratchet.
Ensure the tilt ratchet does not “slip” or ‘chatter” when tilted due to a loose or broken cross bar that will not engage into the teeth. Check the integrity of this mechanism by placing the end of a blade post into the ratchet and seeing how easily it slips when pressure is put on the top of the blade.
Make sure each ratchet can smoothly be placed on the ring without binding. Look into the slot where the ring engages and inspect for raised edges or bent materials.
RoTilt ratchets should rock from side to side while being engaged. The cam should also immediately disengage without “sticking” and there should be no residue or scarring in the slots of the cam where the internal pin locks the RoTilt in place.
Recommendations:
50-4695 Self-Retaining Low Profile Locking RoTilt

## Slide 27: RoTilt ® Ratchets
50-4589
Original RoTilt
High profile
Rotated position held in place by tension on the blade when retracting
Rotated position released when the tension on the blade is released
50-4695
Second Generation RoTilt
Lower profile
Self-retaining clip holds ratchet on ring
Rotated position controlled by thumb piece on side
50-4696
Current Version of RoTilt
Lower profile
Rotated position locked in place by tension on the blade when retracting
Rotated position released by lever on the side

## Slide 28: Replace Competitive Ratchets with Tilt or RoTilt Ratchets

## Slide 29: Replace Old Tilt Ratchets with New Tilt or RoTilt Ratchets
1983
Original Tilt Ratchet® Mechanism – This added an extra feature (can be tilted) to the original regular (straight) ratchet mechanism. Thumb piece engagement is secured by leaf spring mounted at bottom of unit.

## Slide 30: Replace Old Tilt Ratchets with New Tilt or RoTilt Ratchets
1986
Angle of thumb piece was reduced to almost flat to give comfort to the users and also to prevent possible slippage when adjustment is made to the tilted position. Material was also changed to make thumb piece stronger.

## Slide 31: Replace Old Tilt Ratchets with New Tilt or RoTilt Ratchets
1990
Added thickness to the ratchet teeth portion of the casting to eliminate side bending. Made it very strong. Single control feature (tilting and retractor blade in and out movement) was incorporated to achieve both actions by simply depressing the ratchet pawl.

## Slide 32: Replace Old Tilt Ratchets with New Tilt or RoTilt Ratchets
1993
Based on surgeon feedback, they would prefer to control the tilting and the retractor blade movement separately, so the single control feature was removed. This also prevents possible disengagement of the tilted portion if the pawl is accidentally depressed.

## Slide 33: Current Tilt Ratchet
1996
Improved tolerances/fitting of parts. Pivot point of thumb piece was also moved to maximize engagement with ratchet teeth.

## Slide 34: Malleable Blades
Check for possible gouges on blade.
Inspect where the blade meets the post – check for cracking.
Hold Blades up to the light to check for cracks.
Check teeth on the shaft. Make sure they fit into the ratchet properly.
Slide a ratchet down the post to make sure the post is not bent.

## Slide 35: Replace Old or Competitive Blades with New Blades
Original Codman® Blade
Symmetry
Surgical® Blade
Aesculap ®
Blade
Symmetry Surgical ® Blade
Wider space between post notches
Angles instead of curves on Kelly blades
Narrower space between post notches for finer adjustments

## Slide 36: Questions